HMG Series EO Sterilizer

Computer Control System（CE）

HMG系列环氧乙烷灭菌柜

电脑控制系统(CE)

OPERATION MANUAL

操作手册

Jiangyin Huaqing Machinery Co., Ltd

江 阴 市 华 青 机 械 有 限 公 司

Item  目录

- System brief introduction     系统简介

- On-time data manage        实时数据管理

  - Parameter setting            参数设定

  - Product list input             物品清单的输入

  - Work process                工作流程

  - Select work process to start    选择开始步骤

  - Operate method select        工作方式选择

  - On-time data save            实时数据存盘

  - On-time data tend drawing    实时数据曲线显示

  - Alarm                      报警

  - Open/close door              开/关门

  - Heating                     线性加温

  - Air moval                   线性抽真空

  - Keep pressure               保压

  - EO injection                 加药

  - Sheet print                  报表打印

- History data manage          历史数据处理

  - Data backup                 数据备份

  - Data resume                 数据恢复

  - Print history data records      打印历史数据报表

  - Show history tend drawing     显示历史趋势图

- Statement                   声明

1.    System brief introduction    系统简介

EO sterilizer computer control system uses top grade Pentium industry control computer by the military standard, it could collect the data of humidity, temperature and pressure, and condition them, and save the data, and print the data; and the computer could collect and condition the parameter of the position of the door; and the computer could control the parameters of the alarm, and heating switch, and door switch, and so on.

The system uses 32 bites on-time operation system of America Microsoft company, the operation plate is windows 98, it could measure and control on-time with other up-grade programs, the system could run smooth and credible. Meanwhile, windows operate media could make the communion more easily; the operators could operate the system more easily. Mouse mostly does the operations, the keyboard is used only when parameters input.

- On-time data manage         实时数据处理

When the system is opened, the user should input the key words to prove the validity. If the user has selected on-time data manage, then the system run the program.

  - Parameter setting             参数设定

To the parameter setting interface, the user should input the work parameters for sterilizer, such as: the weight of EO input, exposure time, pressure value when the air moval finished, air moval time, chamber temperature value, chamber heating time, the temperature of water tin during chamber heating, the temperature of water tin during exposure time, the temperature of gassing equipment, time interval for data saving and printing, the pressure value at the end of EO injection, etc.

The user also could set the alarm values of temperatures, pressure and humidity.

If the user has something relative to sterilize should input into the final record, the system provides three plant position, the user could input them into the plants.

If some of the upper parameters and alarm values are wrong, the system could awake the user to change them. If all the parameters are correct, then keystroke the “ok” button, the system could save all the setting values. When a second cycle begin, the system could find all the settings of last time, if no change in this cycle, keystroke button to next interface.

The chamber number, order number, sterilize lot must be input, or the program wouldn’t go on.

ATTENTION: pay attention to the sterilize lot ( use the date is better), it could be changed during follow steps of the cycle.

  - Product list input                物品清单的输入

The product list will appear on the final record. The product list input has also been done in the parameter setting interface. The operate could input the product name, spec/type, amount and remark by keyboard, and then keystroke “add a product” button, to add the above information into product group. The system only could add six products into the group. If the user has some different requirements, the manufacture could do some changes to meet them.

  - Work process                 工作流程

    - Open the door of the chamber

    - Close the door of the chamber

    - Chamber heating precondition, and pre-humidity and heat the gassing equipment

    - Air moval, keep pressure

    - EO injection, sterilize

    - Air change

    - Open the door, unload product

    - Records print

  - Select work process to start    选择开始步骤

In the normal operate, operator may start the program from the door operate, and could start the program at any step of the flow chart. When the start step was selected, keystroke the button before the step in the flow chart, and then keystroke the “ok” button. If the program was started after the door closing, the “work start” button will appear only when the doors are closed, and then the program will record data and control the chamber.

When the work started, operator could watch the flow chart at any time, but keystroke the button of the chart, it will not act any more.

  - Operate method select        工作方式选择

The system provides two work methods, work method 1 is the normal operate grogram, work method 2 is designed for validation. It will work as the operator keystroke the relevant button.

  - On-time data save            实时数据存盘

It is designed for the further, to save the on-time data into local HD, the data will be saved, and it could be read in the further. The time interval for saving and printing could be set by operator in the relative interface, the saved file will be named as: “Q” + sterilize lot + “.dat”. The user could find the file out to read the relative parameters.

The saved file records the follow information: work method, product list, sterilize parameter, and sterilize information and on-time data.

The operator could delete the file, but once it was deleted, it couldn’t be find any more.

  - On-time data tend drawing     实时数据曲线显示

The on-time tend shower used the dynamic reference frame, that mean, for one hour, the time reference frame is one hour, for two hour, the time reference frame is two hour, and so on. It could assure that all the data tend could appear in the same screen.

The operator could use the two arrow keys (“↑” “↓”) to see all the parameter on-time tend.

  - Alarm                     报警

The alarm system includes “gassing equipment temperature too high”, “chamber temperature too high”, “chamber temperature too low”, “entrance door open”, “entrance door close”, “exit door open”, “exit door close”, “humidity too high”, “humidity too low”, “chamber pressure too high”, “leakage alarm”, etc.

  - Open/close door            开/关门

The operator could stroke the open/close door sign in computer or turn the switch to open/close door. When use the mouse to keystroke “door open”, alarm rings, and 10 minutes later, the lock unlocked, and 10 minutes later, the door opened. When use the mouse to keystroke the “door close”, alarm rings, and 10 minutes later, the door closed, when the door reaches the position, the lock locked.

When the system starts, it could test the position automatically, and then do the corresponding operation per the doors position. The operation wouldn’t go on with even one door not reach the closing position. So now close the door, the operation will go on to next step.

  - Heating                   线性加温

- If the setting preheat time is “0”, the system will run as full heat;

- If the setting preheat time isn’t “0”, the system will run linear heat in the setting time. If it couldn’t reach the setting temperature when time is over, the system will continue to heat until reach the setting temperature.

- After full heat or linear heat, the system turns into keep temperature status. The keep temperature time should be set in “parameter setting” interface.

  - Air moval                 线性抽真空

- If the setting air moval time is “0”, the system will run as full air moval;

- If the setting air moval time isn’t “0”, the system will run linear heat in the setting time. If it couldn’t reach the vacuum value when time is over, the system will continue to run air moval until reach the setting vacuum value.

  - Keep pressure              保压

When air moval finished, the system will run “keep pressure”. The keep pressure time is set 10 minutes by the system. If the chamber pressure change too obviously (more than 1.8Kpa), the alarm will run, and “leakage alarm” will turn to red and flash. The operator should stop the cycle, and to check out the leakage position. When all is finished, the operator could restart.

  - EO injection               加药

EO injection is started from the sterilize step, at that time the system will open the injection pipe automatically, and the gassing equipment will reach the setting temperature. The step could be controlled by computer or manually. At the beginning of injection, “injection finished” button will appear in the interface. When the chamber pressure reaches setting pressure value, that mean the injection step is finished, the injection pipe will closed automatically, and the “injection finished” button will disappear automatically. On the other hand, the operator may keystroke the “injection finished ” button when the injection is finished, the system will close the pipe, and send a signal to system, and the button will disappear.

  - Sheet print               报表打印

When a cycle finished, the system will print the “EO-sterilization record per load”. If the user wants to print other information, such as parameter records, full parameter tends, and the parameter tends in the first sterilize hour, the operator should keystroke the print list by mouse at first, the system will show a “print select” window, the operator should select to print “parameter record”, “parameter tend”, “an hour parameter tend”, “EO-sterilization record per load”.

Attention: before print records, there should be enough paper in the printer. The system will show a “paper lack” window when the paper used out. The operator should double stroke the printer sign in the right corner, and then cancel all the unprinted files.

- History data manage        历史数据处理

When reach the first interface, user should input the system keys. If the user chooses history data mange, the system will run the program.

  - Data backup               数据备份

Data backup is used to backup sterilization files from local disk to backup disk. The files could be used when necessary. Users should choose backup disks carefully.

When some files need to backup, stroke the “backup data file” at first, double stroke the files in “files for condition” group, the files will list in the “selected files” automatically, when all the files that should backup have be selected, stroke the “ok” button, the system will backup the files automatically.

If the system shows errors window during backup operation, it is mostly caused by leak-connected to backup disk or disk errors. The user should check them at once and then re-operate.

  - Data resume               数据恢复

Data resume is to resume the sterilization data files from backup disk to local disk, to print former records or show former history tend. Users should choose backup disks carefully.

When some files need to resume, stroke the “resume data file” at first, double stroke the files in “files for condition” group, the files will list in the “selected files” automatically, when all the files that should resume have be selected, stroke the “ok” button, the system will resume the files automatically.

If the system shows errors window during resume operation, it is mostly caused by leak-connected to backup disk or disk errors. The user should check them at once and then re-operate.

  - Print history data records      打印历史数据报表

The system provides the method to print history data includes EO-sterilization record per load, EO-sterilization parameter records, sterilization parameter tends.

Choose the file that need to print in “files for condition” group, and then stroke the “print history data” button, the system will show a “print select” window, user could select one or more subject to print, then stroke “ok” button, the selected subjects will be printed then.

Attention: before print records, there should be enough paper in the printer. The system will show a “paper lack” window when the paper used out. The operator should double stroke the printer sign in the right corner, and then cancel all the unprinted files.

  - Show history tend drawing      显示历史趋势图

Users may watch the history tends of each parameter, and then decide whether it needs to print.

At first to select the file that you wanted to see it’s history tend, and then stroke the “show history tend” button. The system will show a history tend window. The operator could use the two arrow keys (“↑” “↓”) to see all the parameter history tends.

- Statement                    声明

  - EO sterilization chamber computer control system must be run in the rooms with air-condition, and connected with earth well.

  - All the settings in the system mustn’t be changed, and no games are permitted to setup in the system.

  - The lines of pressure and humidity sensors mustn’t be changed without permission, and so do the input and output control lines.

  - Any accident caused by the above conditions, user should be response for it.